## Slide 1: REIS-Extended Runaway Electron Imaging Spectroscopy
Giorgio Ghillardi, MArco passeri, valerio piergotti, giuliano rocchi, benedetto tilia, Alessandro sibio, basilio esposito, Daniele carnevale, chiara monti
Fusion Technology Division - ENEA
8th Runaway Electron Modelling (REM) meeting
Jan 2020 - GOTHENBURG

## Slide 2: Presentation Overview
1-REIS project
  Overview
2-REIS-E system
  Components
  Calibration
3-Acquisition tool
  LabView program
  Acquisition phases
4-Synchrofit
  Model
  GUI
5-FTU data and preliminary analysis
  Raw data
  Signal focus
  RE Energy vs shot conditions

## Slide 3: 1 REIS Project
The aim of the REIS project is the design, development and construction of a portable Runaway Electron Imaging Spectroscopy (REIS) system for use in medium size tokamaks
The REIS detects the visible and infrared synchrotron radiation spectra emitted by runaway electrons (RE) during the various phases of a plasma discharges including the runaway plateaus following disruption events.
Notes: portable system for mid size tokamaks

## Slide 4: 1 REIS Project
The REIS-Extended is an upgrade of the first system in order to cover spectra range from 0.4 to 5.0 µm
REIS-E
REIS

## Slide 5: 2 REIS-E components
2 wide-angle objective lenses coupled to the 2 fibers (one for forward view and one for backword view) and with a CCD camera.
Trifurcated fiber for the 3 different spectrometers for FW view
2 fibers for the forward and backward view
MIR range spectrometer
1500-5000 nm
NIR range spectrometer
900-2500 nm
VIS range spectrometers
For both views
190-1100 nm
Fast JAI CCD camera. 120 frames per second

## Slide 6: 2 Calibration
Scheme of the FTU midplane cross section, showing the location of REIS in port 6 (P6). The camera and spectrometers field of view (FOV) are shaded, respectively, in gray and brown.
The backward view is taken for cross checks.
Forward
Backward
Protective BK7 glass window 8mm
CaF2 wide angle lenses
REIS-E Head
REIS tube
VIS
NIR
MIR
Trifurcated fiber
VIS
Camera control unit
Layout scheme. The calibration was performed considering all the components combined in the system (also the tube).

## Slide 7: 2 Calibration
Lamp
REIS-E
Diffusor
Light
Calibration in lab with a Black body Tungsten source and a diffuse reflective surface .
M Lamp radiation in W/
L counts per wavelength from the spectrometer
𝜽 angle between lamp and REIS-E
𝜌 reflectance from the diffuser

## Slide 8: 2 Calibration
The calibrations curves obtained with diffusor reflectance and lamp irradiance, as mentioned before. Below a detail for NIR curve.
Raw data

## Slide 9: 3 Acquisition Tool
The acquisition system is synchronized with the main common gate of the tokamak. Trigger time phasing for acquisition in external hardware triggering is represented in figure for FTU case.
The integration time window is 15 ms, but the triggers are 20ms distant in order to allow a correct data saving

## Slide 10: 4 Analysis Tool: Synchrofit
Synchrofit is a python software made in ENEA for fitting Synchrotron radiation data from different tokamaks
Features:
    Fitting data with monoenergetic energy distribution (integration with different distribution model planned)
    Friendly User Interface
    Multiple tokamaks configuration
    Signal and energy threshold to check signal presence
    Single lines emission rejection
    Savitzky-Golay filter on calibrated data
    Total temporal evolution of shot parameters

## Slide 11: Fitting Model
Acquired spectra are fitted with RE Monoenergetic distribution model.
The total brightness seen by the diagnostic (ref.1)
Where
is the pitch angle defined as
is the Lorentz factor, related to the kinetic energy of the particle by
4 Analysis Tool: Synchrofit
1. G. Bekefi, Radiation Processes in Plasmas, edited by S. C. Brown (Wiley & Sons, 1966)

## Slide 12
The single particle emission spectra is obtained by the formula
Where is the critical emission wavelength expressed by
Fitting Model
4 Analysis Tool: Synchrofit

## Slide 13: 4 Analysis Tool: Synchrofit
is the Bessel modified function of second kind. The integral is solved using the following identity involving ipergeometric functions (ref 2)
Validity of the model:
The estimate for the brightness is valid for highly relativistic electrons
Runaways scattering off cold background ions negligible
Cylindrical approximation is valid only for small pitch angles and does not take into account the effects of magnetic field line curvature and curvature drift
2. http://functions.wolfram.com/Bessel-TypeFunctions/BesselK

## Slide 14
4 Synchrofit GUI
Here an example of the GUI view.
In detail this panel shows the calibrated signal from every spectrometer (the empty one is waiting the MIR calibration).
In the first 2 graphs it can be seen the signal difference from the backward and the forward view during the RE beam of shot 43654

## Slide 15: 5 FTU data and preliminary analysis
FTU autumn data campaign started the 24 October 2019 and ended the 29 November (first programmed shot 43321)
8 days dedicated to RE programs
The MIR spectrometer was added from shot 43647, with a manual acquisition start
The CCD camera acquisition system was only integrated during the final phase of the campaign
6 RE beam recorded

## Slide 16
5 RAW data
Example of the raw signal, including MIR spectrometer for shot 43654 at 630 ms.
Shot 43654
In the top figure a view of the discharge: plasma current and neutron/gamma detectors signal.
The red dashed line represent the signal reported here at time 630ms
The gap recorded around 2700 nm is related to the BK7 glass window.
nm

## Slide 17
Example from shot 43651
nm
counts
5 RAW data
In some cases saturation occurs in VIS and NIR spectrometers
Need to decrease time acquisition window during RE beam

## Slide 18: 5 Spectra Fit
Fitting parameters
  Pitch angle 0.1
  Wavelength window
    VIS 500-800nm
    NIR 1100 – 2300 nm
  Savitzky-Golay filter
    Filter order: 3
    Window length: 7
Example of spectra fit from shot 43654 during RE beam
The VIS range is not well described by the curve. New fits with modeled distributions functions are needed.

## Slide 19: 5 Spectra Fit
shot 43654
nm
Ph/s/m^2/nm/str
The fit in separated spectra show that the VIS range is described by higher energy around ~30 MeV while the NIR range is described by energy around 21/23 MeV.
This behavior is seen also for other shots.
Ph/s/m^2/nm/str
nm
shot 43654
NIR
FW

## Slide 20
310 ms
330 ms
350 ms
370 ms
Images from shot 43654

## Slide 21
370 ms
Time (ms)
350 ms
Images from shot 43654
FW view with RE beam
True runaway signal corresponds to blue line (i.e. backward signal subtracted off)

## Slide 22
shot 43654
5 Spectra Fit
The monoenergetic distribution doesn’t describe properly the VIS range during the RE beam.
Need to consider different energetic distribution function instead of monoenergetic model.
shot 43654
With BW subtraction

## Slide 23
5 Spectra Fit
BW subtraction improves quality of runaway spectrum fitting during ohmic phase of discharge.
shot 43654
shot 43654
With BW subtraction

## Slide 24
Time (ms)
5 Energy Focus
The signal got saturated in shot 43654 in shadowed zones.
Fitting procedure still to be properly corrected in saturated zones
BW subtraction only affects energy evaluation in ohmic phase and final phase of RE beam
v

## Slide 25
In the second panel the shape of the neutron BF3 scintillator shows a good agreement with the spectra integral.
The neutrons are produced through photo-neutron reactions when the energy of the photons generated by the runaways is > about 10 MeV.
Time (ms)
5 Neutron Signal

## Slide 26: 5 Number of RE in RE BEAM
Shot #43654
where is the plasma current and is the cross-sectional area of the runaway beam
RE beam
The estimated n_RE from brightness is clearly increasing after the RE beam starts
The number of RE is decreasing during the RE beam
The number of runaway is related to the plasma current by the relation

## Slide 27: 5 DATA Analysis
6 RE beams recorded:
  43521: Bt = 5.5 T , Ip flat-top ref = 350 kA
  43525: Bt = 5.3 T , Ip flat-top ref = 350 kA
  43539: Bt = 5.3 T , Ip flat-top ref = 353 kA
  43540: Bt = 5.3 T , Ip flat-top ref = 500 kA
  43651: Bt = 5.4 T , Ip flat-top ref = 351 kA
  43654: Bt = 5.5 T , Ip flat-top ref = 352 kA
Focus on the RE beam in shot 43521

## Slide 28: 5 DATA Analysis
6 RE beams recorded:
  43521: Bt = 5.5 T , Ip flat-top ref = 350 kA
  43525: Bt = 5.3 T , Ip flat-top ref = 350 kA
  43539: Bt = 5.3 T , Ip flat-top ref = 353 kA
  43540: Bt = 5.3 T , Ip flat-top ref = 500 kA
  43651: Bt = 5.4 T , Ip flat-top ref = 351 kA
  43654: Bt = 5.5 T , Ip flat-top ref = 352 kA
Focus on the RE beam in shot 43540

## Slide 29: 5 Runaway population in ohmic discharge
4 T, Iref =500kA
RE energy clearly increase during the discharge.
Current flat top shot 43454

## Slide 30: 5 REIS/REIS-E
The first plot show the result obtained from the first REIS system.
RE energies are in alignment with old results: RE beam energy between 20-30MeV for all 6 RE beams.
The REIS-E data will be re-fitted with the MIR range (after MIR spect calibration)
v
Time (ms)

## Slide 31: 5 Error evaluation
Work is in progress to evaluate the uncertainties in the spectral data:
Systematic
  Calibration (~10%)
  Tokamkas variables (Btor, Radius)
Statistical
  Fitting uncertainties
  Window fitting size selection
  Background subtraction
  Counting statistics

## Slide 32: Outcome and future work
Status
REIS-E assembled and tested in FTU runaway discharges (2019)
REIS-E can be installed in 2020 in mid-size tokamaks (e.g. COMPASS, AUG)
Work on-going
Calibration of MIR range
Analysis of FTU runaway electron data in full wavelength range from 0.4 to 5 micron
Error analysis
Upgrade of Synchrofit with non monoenergetic distribution functions
Finalize set-up of CCD camera and perform imaging data analysis